## Slide 1: Contact-Free Heart and Respiration Rate Measurement System Based on Real-Time Human Face Video
Eugene Dobryakov - Team Leader
Quan Nguyen
Roman Melnik
Faculty Advisor - Dr. Essa
Team / Project - 115

## Slide 2: Outline
Executive Summary
Background and Problem Statement
Objective
Technical Approach
Experimental Results
Deliverables
Project Timeline/Gantt Chart
Budget
Professional Awareness
Conclusion

## Slide 3: Executive Summary
The heart is at the center, and the most important organ of the
circulatory system.
Elevated heart rate and respiration rate are physiological signs
associated with COVID-19.
It can take just one individual infected with COVID-19 to spread
the virus to at least two additional people [1].
Heart rate (HR) and respiration rate (RR) can be measured in
real-time through video without ever coming in contact with
another individual.

## Slide 4: Background - Did You Know..
U.S. adults avoid medical care due to concerns regarding COVID-19.
4 in 10
Delay or Avoidance of Medical Care Because of COVID-19–Related Concerns
http://dx.doi.org/10.15585/mmwr.mm6936a4
Increase in mortality rate for COVID-19 patients whose respiratory rate exceeds 22 breaths per minute.
3.2 fold
Admission Respiratory Status Predicts Mortality in COVID 19
https://onlinelibrary.wiley.com/doi/10.1111/irv.12869
50%
COVID-19 Healthcare Coalition
https://c19hcc.org/telehealth/patient-survey- analysis/
Survey participants stated they would delay medical care during COVID-19 pandemic without access to telehealth services.

## Slide 5: Problem Statement
Measuring HR
  Clinical Standard - ECG [2]
  Leading Method For Individuals - Manually Checking Pulse [3]
  On The Rise - Wearables Such as Fitness Trackers
Challenges
  ECG - Factors such as pregnancy and obesity interfere with results [2]
  Manually Checking Pulse - Not Accurate
  Wearables - Measures blood flow when its farthest away from the heart
Opportunity
  Re-evaluate delivery of health care during and subsequent of the pandemic
  Lift and shift system to measure HR and RR in real-time through video

## Slide 6: Objectives
1
Extract region of interest (ROI) with an RGB-D camera
2
Extract the raw RGB signals from the ROI
3
Convert signals to a frequency domain to obtain HR and RR
4
Display HR and RR to the subject

## Slide 7: Technical Approach
Software Architecture Overview

## Slide 8: Feeding Live Video Into the Program
In real time -> live video input
High frame per second to compensate for processing time
Wired connection
Improved security

## Slide 9: Selectin the ROI
Why the forehead?
Good accessibility
Minimum computing area
Not covered by clothes or accessories
Proved to be efficient in previous studies [5]

## Slide 10: Detecting the Forehead
Pretrained AI model using CMake and Dlib libraries
Model detects 68 facial landmarks
Need 4 landmarks: 20, 25, 38 and 44
Mirror 38 and 44 across 20 and 25
Convert input frames to grayscale to process
68 Facial Landmarks

## Slide 11: Using the Green Channel
Easier for the skin to absorb the green color
High signal-to-noise ratio [6]
Better reading => more precise output
https://en.wikipedia.org/wiki/RGB_color_model

## Slide 12: Extracting Raw Signal
Store data in time domain: Extract signals of ROI from input frames, calculate mean ROI Value for the first several seconds and store in a buffer.
The program matches the mean ROI value with the corresponding time.
For best FFT function results, size of time domain should be a value of 2n, where n is the number of sample points (i.e., frames)
Mean ROI value for each frame
Corresponding time for each Mean ROI value

## Slide 13: Conversion Process
2n frames captured over time period Δt
  Frame 0 at t0, frame 2n-1 at t-1
Create an evenly spaced time period list (ESTPL)
  Example: 64 frames captured at 8fps, means ESTPL is [0.000, 0.125, 0.250, …, 7.875]
Detrend values to remove linear trend
Generate a list of incrementing frequencies via NumPy Function
  [0.122890625, 0.24578125, …] (later multiply by 60)
  2n frames makes a 2n list of frequency values
    [7.3734375, 14.746875, …]
    This is the first sign of the HR

## Slide 14: Conversion Process
We use interpolation to estimate unknown values (requires ESTPL)
  Allows us to say this is where the data point would have been had it occurred at the time in question
Interpolated data is normalized before being converted to Frequency Domain.
Interpolation - About ~16 Seconds

## Slide 15: Analyzing the Signals
Transform signals from time domain to frequency domain using Fast Fourier Transform (RFFT)
HR: 0.8 to 2.2 Hertz [7] & RR: 0.18 to 0.5 Hertz [7]
Identify the signal peak and look at the same “frequency” bucket.
  If peak in index 8, then the corresponding values for the HR or RR is in the 8th index of the frequency list.
Display the result

## Slide 16: Experimental Results - HR
HR data validation with Apple Watch HR Monitor application
HR Testing Results
Apple Watch HR Monitor

## Slide 17: Experimental Results - RR
RR data validation - physical count of number of breathes in a 20 second time period
Multiplied by 3 to obtain number of breathes per minute
5 breathes recorded during the 20 second window in this test
5 * 3 = 15 breathes per minute
Program detected RR to be 14 breathes per minute

## Slide 18: Deliverables
Contact-free and real-time system
Portable, lift-and-shift system
Compatible on common OS (Windows, MacOS)

## Slide 19: Deliverables
Final Product with Touchscreen Monitor and Camera

## Slide 20: Project Timeline/Gantt Chart - Fall 2021

## Slide 21: Project Timeline/Gantt Chart - Spring 2022

## Slide 22: Budget
Item | Supplier | Quantity | Unit Price | Total
Intel RealSense Depth Camera D455 | Intel | 1 | $433.94 | $468.88
12 Inch Touchscreen Monitor HDMI 1920x1080 FHD IPS Screen | TOGUARDGO | 1 | $129.99 | $140.39
 |  |  | Total: | $609.27

## Slide 23: Professional Awareness
Important that data from the ROI be kept private
Evolving technology provides room for future advancements
Contributes to well-being of society as a tool against COVID

## Slide 24: Conclusion
Contact-free, real-time HR and RR measurement system with multiple benefits
Positively impacts telehealth assessments.
Helps to detect elevated heart rate and respiratory rate.
Advances the delivery of medical service during and subsequent of the pandemic.
System can be moved from point to point with minimal effort.

## Slide 25: Questions

## Slide 26: References
[1] Nazario, Brunilda. “How Coronavirus Is Transmitted: Here Are All the Ways It Can Spread.” WebMD, WebMD, 18 Aug. 2021, https://www.webmd.com/lung/coronavirus-transmission-overview#1
[2] “Electrocardiogram.” Johns Hopkins Medicine, https://www.hopkinsmedicine.org/health/treatment-tests-and-therapies/electrocardiogram
[3] “Want to Check Your Heart Rate? Here's How.” Harvard Health, 17 Aug. 2021, https://www.health.harvard.edu/heart-health/want-to-check-your-heart-rate-heres-how
[4] Team, Heart and Vascular. “Your Fitness Tracker Isn’t the Best Way to Measure Heart Rate.” Cleveland Clinic, Cleveland Clinic, 18 May 2020, https://health.clevelandclinic.org/your-fitness-tracker-isnt-the-best-way-to-measure-heart-rate/
[5] S. Kwon, J. Kim, D. Lee and K. Park, "ROI analysis for remote photoplethysmography on facial video," 2015 37th Annual International Conference of the IEEE Engineering in Medicine and Biology Society (EMBC), 2015, pp. 4938-4941, doi: 10.1109/EMBC.2015.7319499.
[6] Y. Lin and Y. Lin, "A study of color illumination effect on the SNR of rPPG signals," 2017 39th Annual International Conference of the IEEE Engineering in Medicine and Biology Society (EMBC), 2017, pp. 4301-4304, doi:10.1109/EMBC.2017.8037807.
[7] Shourjya Sanyal and Koushik Kumar Nundy, “Algorithms for monitoring heart rate and respiratory rate from the video of a user’s face,” IEEE Journal of translational engineering in health and medicine, vol. 6, pp. 1–11, 2018. https://www.rouast.com/pdf/rouast2016remote_a.pdf